Memorial Health System

Security and Technical Questionnaire

| Product Name: | PeriWatch Surveillance |
| --- | --- |
| Requestor: |  |
| Vendor Name: | - PeriGen |
| Date: | - 3/3/2021 |

# I. Infrastructure

| Question | Response |
| --- | --- |
| Please provide a visual diagram of the overall process of how this product/technology will be used within Memorial. | See attached |
| Will this application / solution be installed on hardware purchased by MHS or will this be a fully vendor supplied or will it be hosted in the cloud? | All hardware will be purchase by MHS. |
| Does the application require remote access from the outside for support? If so does your company support the use of SecureLink? | Yes – SecureLink is our preferred connection method. |
| Does the application / solution provide for high - availability clustering, load balancing or other DR / BC capabilities? (Please provide diagram or DR/BC architecture.) | Product supports VMware HA and SQL clustering. |
| Please provide a visual diagram that shows the client / server communications and technologies involved with your fully functional product. | See attached |

# II. X86 Architecture

| Question | Response |
| --- | --- |
| What Server Operating Systems are supported? | Server: Windows 2012R2, Windows 2016 SQL: 2014 & 2016 |
| What is the validation cycle for new Server Operating Systems? | There is no set time frame. PeriGen follows market trends and validate new operating systems once customers are adopting them. Practically, PeriGen validated 2 main server OSes, for example, presently Windows 2012R2 and 2016 are still supported. With version 2021.1, Windows 2019 will be validated. With version 2022.1, support for Windows 2012R2 will most likely be dropped. |
| What is the notification process for older Server Operating Systems being retired? | PeriGen provides release note for upcoming versions that include listing of system requirement for the version. |
| What Desktop Operating Systems are supported? | Workstations: Windows 10 |
| What is the validation cycle for new Desktop Operating Systems? | Based on customer base adoption of new operating systems. There is no set cadence otherwise. |
| What is the notification process for older Desktop Operating Systems being retired? | Typically, inline with Microsoft announcing OS EOL |
| Does your application (and all components) support virtualization on VMWare? | Virtualization is supported for Application and Database Servers. |
| If no, what components need to reside on physical servers? | Serial Servers are normally used to capture serial communication from fetal monitors. |
| What versions of VMWare are supported? | VMware (6.X and above) and Hyper-V is currently supported |
| Does your application (and all components) support running in Citrix? | Yes. |
| If no, what components are not supported? |  |
| What versions of Citrix are supported? | Citrix XenApp 7.15 LTSR (Long Term Service Release) and above installed on Windows Server 2012 R2 or Windows Server 2016. Additional versions may work but will require testing. |
| Does your system rely on a user profile? | Not on Windows user profiles. We do depend on AD user for authentication to the application(s) |
| Does your application support Citrix universal printing? | No |
| Assuming that a given Citrix server is dedicated to your application / solution, how many user sessions are supported on one server and what are the specifications? | To be determined as part of project implementation sizing. However, a rule of thumb each session use ~150MB of memory and ~30 concurrent sessions per server can be targeted. See overall specifications attached. |
| If your application requires a physical server, can it run on a blade server? | VM server are supported. |
| If no, please list the reason why (ie physical dongle, proprietary card etc...) |  |
| Does the application / solution support the Symantec Enterprise AV solution? | Yes. Certain directories will need to be excluded. |
| Please list any specific exclusion requirements. | See PeriWatch Surveillance AV Exclusion Guide |
| Does your application / solution support being a domain member? | Yes |
| If no, please detail how enterprise management, system patches, and user account administration takes place. |  |
| If your application / solution or any component is web based, please provide ALL supported Web Browsers and versions. | MS IE 11 is supported but not recommended due to very slow Javascript processing. The original MS Edge is supported but not recommended. The latest Edge Chromium is recommended (89+). Google Chrome will provide best performance. |
| Identify all memory & CPU requirements for each component of your solution both at the server and client levels. | See attached System Requirements document |

# III. Storage Architecture

| Question | Response |
| --- | --- |
| Does the application / solution support TSM Backup solution? | Yes |
| If no, please identify supported backup solution. |  |
| Does your application / solution have any special back up requirements? (ie downtime required for proper back up because of open file limitations) | No |
| Identify any performance limitations of your applications related to Industry Standard IOPs measurement. | None on modern SAN/storage infrastructures. |
| Does your application / solution support Fiber SAN disks? | The PeriGen application does not directly interacts with storage. Assuming this is part of an enterprise level VMware infrastructure, SAN infrastructures are supported. |
| If no, please identify the raid configuration for local disk. |  |
| Does your application / solution have any other limiting storage requirements? | No |
| What is the estimated storage growth per year for a health system Memorial's size? | Rule of thumb is <10MB per delivery. So if across the enterprise you have 3000 deliveries per year = < 30GB |

# IV. Mid Range Architecture

| Question | Response |
| --- | --- |
| Please identify all supported OS's. (ie, AIX, Cento OS, RedHat Linux, etc.) | Windows 2012R2, 2016 and 2019 starting version 2021.1 expected to be released in July 2021. |
| Please identify all supported hardware platforms. (ie IBM P-Series, HP Itanium) | PeriGen assume a VM mode of deployment that complies with the system specifications provided. Hardware brand or model should not be a factor providing the resources specified are met or exceeded. |

# V. Network Considerations

| Question | Response |
| --- | --- |
| Does your application / solution have VPN requirements? | Not for application use. This is a 100% customer hosted platform. |
| Does your application / solution have wireless requirements? | No |
| Does your application / solution have modem requirements? | No |
| Does your application / solution have voice/phone hardware requirements? | No |
| Does your application / solution have voice/phone connectivity requirements? | No |
| Does your application / solution have an available port count requirement? | No |
| Does your application /solution have power requirements? (ie plug type, watts or # of power supplies) | No |
| Identify any bandwidth requirements that your application / solution may have. | LAN speeds with <20ms latency |
| Identify any firewall ports that need to be open for your application / solution to function, if any. | Not required for LAN deployment. The list of communication ports required for the application models is listed in the last section of the attached System Requirements document. |

# VI. Technical Services Considerations

| Question | Response |
| --- | --- |
| Does your application / solution have file share requirements? | Yes. Typically, this is used for archiving and Cold-Feed (pdf export) of charts (is the customer is purchasing this module). This is typically discussed during the pre-install phase. |
| Please provide a yearly estimated growth. | # of annual births x 10MB per year |
| For Touchworks related applications – Does your application security model depend on the Touchworks application security permissions/definitions or is it dependent on an external source for authentication or its own application security model? (Do independent user accounts need to be set up in your application or in Active Directory for your application to work?) | NA |
| For Touchworks related applications - If Active Directory authentication is required for authorization do you support ADFS? (Note: We have 3 different organizations with different AD structures so each organization would have a different Active Directory from one another to source to) | NA |

# VII. Interface Considerations

| Question | Response |
| --- | --- |
| Does your application / solution have interface requirements? |  |
| - ADT | Yes |
| - Orders | No |
| - Results | Yes |
| - Other (please specify) |  |

# VIII. Customer Support Services Considerations

| Question | Response |
| --- | --- |
| Does your application / solution require any new equipment such as PC's, monitors, scanners, printers etc? (Please provide spec documents) | No. The assumption is that existing workstations can be used. See System Requirements document |
| Does your application / solution require a unique username/password? | For support and proper application configuration, we expect a domain account and a service account. |
| Does your application / solution require special characters or uppercase? | No. The assumption is that AD/LDAP user authentication will be used. |
| Does your application / solution support the use of the Cisco anyconnect 802.1client? | This was never tried. For support, PeriGen uses SecureLink to access the application server(s) |

# IX. Security and Privacy

| Question | Response |
| --- | --- |
| Data Protection Are your data centers under local compliance requirements? If so, which ones? | NA. All server are hosted by the customer. |
| Do you comply with all applicable laws/regulations, including but not limited to the ones below? |  |
| - PCI regulations? |  |
| - HIPAA Regulations? | Yes |
| - HiTECH, Omnibus, other? | NA |
| Please attach SOC1, SOC2, or other risk assessment documentation. | NA |
| Data Segregation How do you separate Memorial data from other customer's data? | NA – All data is stored on customer provided servers. |
| Data-at-Rest Protection Where do you store Memorial's data? | All data is stored in MS SQL databases. The application supported MS SQL TDE for full DB encrypt |
| How strong is your encryption and data integrity? | See MS SQL TDE specifications. Customer would provide the encryption certificates. |
| What kind of authentication and access control procedures are in place? | SQL user authentication |
| Data-in-Motion Protection How do you get data from Memorial to you? (Data encrypted? VPN? Please specify details) | All data in transit is on the internal network only. Data in transit in secured with TLS encryption over https where the customer is providing and managing the certificates used for communication. |
| How do you transfer data from one place to another? (Data encrypted? VPN? Please specify details) | NA |
| What are your data leak prevention capabilities? (if applicable) | NA |
| Can any third party (your service providers) access MHS data, and if so, how? | No |
| What visibility will we have? | NA |
| Please provide a copy of your security practices. | See attached |
| How do your organization ensure that all MHS data is erased at the end of service? | All data is stored on customer provided and managed server. PeriGen does no delete or erase any data stored on customer servers. |
| Vulnerability Management How can you provide evidence of your vulnerability management program? | Veracode scan output |
| How often do you scan for vulnerabilities on your network and applications? | We use Veracode to scan the PeriWatch platform with every version of the product. |
| Can MHS conduct an external vulnerability assessment on your network, and if so, how? | MSH can conduct an assessment of the product installed at CHS. |
| Please list any URLs or Websites that would be associated with this application. | NA |
| Can MHS do a basic Tenable vulnerability scan against any website that is associated with this application? | No external websites are used by the application. |
| Identity Management Are you Active Directory aware, and if so, how? | Yes. User authentications used in the application are AD/LDAP based. |
| Please provide a pictorial of the architecture of integration. | AD users that are part of agreed upon groups are allowed access. At system configuration, users are then associated to authorization groups within the application that provided the more granular access rights (which screens, which specific permissions etc.) depending on customer needs. Authentication is always done against AD/LDAP. |
| Do you keep your own user accounts? | Only the names |
| How do you secure user IDs and access credentials? (please provide details) | Credentials are no saved. |
| How do you handle user churns (e.g., provision and de-provision accounts)? | Since authentication is always done again AD/LDAP, disabled users will no gain access to the app. It is |
| Can you support federation, and if so, which standards? | Single or Multiple distinct domains are supported with the surveillance product. |
| Physical and Personnel Security Is there restricted and monitored access to critical assets 24x7? | NA – Hosted at/by MHS |
| If dedicated infrastructure is desired, can you ensure it is isolated? How often do you scan for vulnerabilities on your network and applications? |  |
| Do you perform background checks on all relevant personnel? How extensive? |  |
| Do you document employee access to customer data? |  |
| Availability How many nines do you guarantee in the SLA? | The application is fully hosted at/by MHS. PeriGen provides 24/7 support services. The product has been in the market for more than 20 years with very high availably record. Since the application is deployed on customer provided hardware/servers/network, SLA can be hard to guarantee. The application typically runs on VMware infrastructures and SQL clusters which provide excellent levels of high availability. |
| What availability measures do you employ to guard against threats and errors? | VMware infrastructures and SQL clusters |
| Do you use multiple ISPs? | NA |
| Do you have DDoS protection, and if so, how? | NA |
| Can you provide availability historical data? | NA |
| What is your downtime plan (e.g., service upgrade, patch, etc.)? | Customer driven. |
| What is your peak load, and do you have enough capacity for such a load? | NA – Hosted at/by MHS |
| Incident response What is your procedure for handling a data breach? | PeriGen does not directly handle MHS data. The system is installed and managed by MHS with support services from PeriGen. However, PeriGen does have incidence response procedures regarding any potential data breach. See the attached Security & Privacy policy. |
| What timeframe does notification occur? | Severity based. High severity items trigger notification within 24 hours. |
| In what format do notifications go out, who are they sent to, and what info do they contain? | Email to the customer primary contact |
| Ensure that the vendor’s incident response procedures do not violate our own incident response requirements. | Unclear |
| Privacy How do you ensure that critical data (e.g., payment card number) is properly masked and that only authorized individuals have access to the entirety of the data? | NA |
| How you protect digital identities and credentials and use them in cloud applications? Pictorial? | NA |
| What data do you collect about me (logs, etc.)? | Basic patient info included in ADT messages (name, MRN, Account #, DOB…) + clinical data from fetal monitors (fetal tracings and maternal vitals) + OB results charted in the EMR and interfaced over HL7 (Cervical exams, membrane status, GA/EDD, GPTPAL, fetal heart analysis, annotation etc.) |
| - How is it stored? | In SQL db as discrete values |
| - How is the data used? | Used to display the information to clinicians in near realtime |
| - How long will it be stored? | No limit |
| - How is it destroyed? | Never destroyed. This information is required for long term archiving |
| Under what conditions might third parties, including government agencies, have access to MHS data? | None |
| Can you guarantee that third-party access to shared logs and resources won’t reveal critical information about my organization? | Yes. MHS data is only available within the MHS network. |
| Patching How often do you apply Security Patches? | Customer driven |
| OS Updates? | Customer driven |
| Is there downtime associated with either of these? | It depends on the patch. Most customers have a regular monthly patch window when patches are applied and servers restarted. |
| Please provide documentation of the process. |  |

# X. Compliance

| Question | Response |
| --- | --- |
| Business Continuity and Disaster Recovery Do you have any DR and BC planning documents, and if so, can we review them? | This is discussed with each customer during preinstall activities. Since the SQL DB is provided/managed by the customer, we have to coordinate/tailor the backup plans to customer policies and infrastructure |
| Can we do a BC audit? | NA |
| Where are your recovery data centers located? | NA |
| What service-level guarantee can you offer under DR conditions? | NA |
| Logs and Audit Trails Can you accommodate timely forensic investigation (e.g., eDiscovery)? | NA |
| Can we agree on provisions in the SLA for investigation? | NA |
| What would we have access to? How? | All hosted at MHS |
| How long do you keep logs and audit trails? Can you keep them as long as we desire? | No limit. All logs are stored I the SQL DB hosted at MHS and dedicated to this application. |
| Can we have dedicated storage of logs and audit trails, and if so, how? | No since logs are stored in the database. |
| Show evidence of tamper-proofing for logs and audit trails | All managed through access control in SQL |
| End-of-Service Support Specify what the cloud vendor will deliver at the end-of-service period. | NA |
| Will data be packaged and delivered back to me? If so, in what format? | NA |
| How soon will I have all my data back? | NA |
| Will any remaining copies of data be erased completely from your network? If so, how soon will it happen? | NA |
| Specify any fees that may incur at the end of the service. | NA |