Date/Time:      August 28, 2019

Vendor:            PeriGen

Address:         100 Regency Forest, Suite 200

Cary, NC 27518

Vendor Reps:   Zach Schooler

Client Reps:

**** In the Observations/Comments column, the grayed-out text suggests what to look for – you may write over this.

| # | Criteria | Observations / Comments |
| --- | --- | --- |
| Services |  |  |
| 1 | Service offerings | Provide continuous near real-time fetal/maternal surveillance for OB department. |
| 2 | SLA | PeriGen provides 24x7x365 support services. The PeriWatch System is developed to be used continuously 24x7. Actual up-time largely depends on the architecture of the system (ex: use of cluster DB, VMware HA etc.) |
| 3 | Segregation of services | All services are provided based from out HQ office in Cary NC. |
| Infrastructure |  |  |
| 4 | Location of data centers | Server (virtual) are to be provided and hosted by the customer. |
| 5 | Tier ratings of data centers | NA. servers are hosted and provided by the customer. |
| 6 | Network design | Hosted on the customer hospital network. |
| 7 | Outage history | NA assuming the question refers to vendor/cloud infrastructures |
| 8 | Uptime | NA assuming the question refers to vendor/cloud infrastructures |
| Application and End User Access |  |  |
| 9 | User interface | Url access, Encryption protocols and required certificates |
| 10 | Access Requirements | Access inside/outside of the GBMC network? |
| 11 | Other |  |
| Monitoring |  |  |
| 11 | Service monitoring (client) | Since the application is meant o be a 24x7 system, any errors or service interruptions are displayed/prompted to the user on the main application client. |
| 12 | Monitoring tools | There are no specific monitoring tools. Extra monitoring on server services can be done using standard Windows notification tools. |
| 13 | Notification process | Typically, the customer notifies PeriGen of a problem OR support detects a problem during routine maintenance/monitoring. In the latter case, support will notify to primary contact of the problem. |
| Recovery Capabilities |  |  |
| 14 | BC/DR Plans | SQL backup agents have been established with the customer and are in place currently |
| 15 | Recovery protocols | Downtime procedures are in place by the clinical team to go to paper in case of failure. Tracing data it always cached at the DAS level and re-synchronized once server/network is reconnected. In case of full system restore from backup, ADT transactions will need to be re-processed, manual data entered for the duration of downtime (depending on customer guidelines) and cached tracings re-associated. |
| 16 | Notification process | Being an onsite system, the customer will be aware of the whole process. |
| Data Protection |  |  |
| 17 | Encryption protocols | All data resides only on customer secure network. On the latest version of the software (2019.X), all data in transit is encrypted (SHA256) using certificates provided by the customer. |
| 18 | Access control | Audit logs are built-into the application and are stored in the SQL DB. It is the customer responsibility to monitor access using the available reports. |
| 19 | Monitoring tools | Built-in reports (via MS Reporting Services) |
| 20 | Notification process | How vendor notifies customers of potential breach? |
| 21 | Data Center Certification | NA. Hosted at customer data centers |
| Workstations and Devices |  |  |
| 22 | Operating System | The latest version of the system (2019.1.02) support Windows 2012/2016 servers and Windows 10 clients |
| 23 | Antivirus | Yes |
| 24 | Network Access | Only on hospital network |
| 25 | Network Sharing | Not required |

| Other |  |  |
| --- | --- | --- |
|  | Fill in |  |
|  | Fill in |  |
| General Impressions |  |  |
| Of the site, of the vendor’s general suitability to meet client requirements. |  |  |

Data Center Tier Ratings

| Tier 1 | Non-redundant components – only what is needed for network links, electric circuits, and HVAC equipment. Single delivery path for power and cooling to servers. |
| --- | --- |
| Tier 2 | Redundant components (aka N+1 where N = needed equipment + 1 extra). Allows for loss of a component without affecting operations. Still a single delivery path. |
| Tier 3 | Redundant components (N+1 or N+2) with multiple network links, power feeds, power gear, and HVAC gear. Dual delivery paths - one active, one passive (A&B). Site is concurrently maintainable (components can be taken offline without affecting operations). Servers and gear can be dual corded. |
| Tier 4 | Fault tolerant, i.e. all components are fully redundant (N+N for network, power, HVAC). Site is concurrently maintainable. All equipment is dual corded and served by dual active delivery paths. Few Tier 4 sites exist. |

Note: A site may have characteristics of more than one tier. It is assigned a particular tier rating only if it meets all of the criteria for that tier.

Review (to be completed by GBMC IT Team)

Approval:

Review Summary: